Operations/Work/Job Activity covered by this assessment: Macro Hardness Tester (Rockwell)

Required Personal Protective Equipment (PPE) : Safety glasses must be worn at times.; Long and loose hair must be contained.. Appropriate footwear with substantial uppers must be worn.; Close fitting/protective clothing to cover arms and legs must be worn.. Respiratory protection devices required.; Oil free leather gloves and spats must be worn.. Hearing protection must be used when using this machine..

- Identify the hardness scale required for the sample under testing.(HRA,HRB,HRC,HR15N & HR15Y)(A)Fig.1 - Turn the knob to Rockwell position for scales HRA,HRB & HRC and Superficial position for scales HR15N & HR15Y (B) (Fig.2 a,b) - Select the load based on the selected scale by rotating the knob.(C) (Table 1) - Mount the suitable indentor based on selected scale.(D)Refer table 1. - Clean the test-block with tissue paper and place it on the special anvil or work table of the machine.(E) - Ensure the loading lever is pulled forward as shown in fig (Fig.3) - Rotate the capston wheel slowly to raise the test specimen till it comes in contact with the indentor tip.(F) - Ensure the corresponding scale is set at zero.(A) - Push back the loading lever as shown in fig.(Fig.4) - After time lapse of 5 seconds record the reading shown on the scale. - Repeat the steps shown in Sr.no 5 to 8 three times and record the readings. - From the readings obtained from sr.no 11, record the deviation from the standard value. - Now the specimen is kept on the anvil and steps 5 to 8 are repeated. - After recording the values, pull forward the loading lever to the unloading position.(Fig.3) (Fig.2 a) (Fig.2 b) SCALE; LOAD; INDENTOR; DIAL SCALE. HRA; 60 Kg; DIAMOND; BLACK. HRB; 100 Kg; 1/16” BALL; RED. HRC; 150 Kg; DIAMOND; BLACK. HR15Y; 15 Kg; ½” BALL; N/T. HR15N; 15 Kg; DIAMOND; N/T. (Fig.3) (Fig.4) (Table.1)